May 23, 2001	Al Boehnlein  Confidential

## GDC Alignment procedure, Rev. 2

# Tools Needed

Allen Wrenches

Tape measure

Optical Fiber Power Meter Level

O’Scope

Hi8 Sony Camcorder and monitor

GDC Alignment Procedure

- Remove M7 to enable better access, be careful not to bump M3.

- Adjust beam to be straight and level and centered through input port and in the center of M1

- Adjust M1 to put beam in center of M2

- Adjust M2 to place beam in corner of M3, as shown in photo 3.

- Adjust M3 to place beam on center of vertical axis of grating, at spot 1.  If necessary, use template cut out from last page of this document.

- Adjust grating to center 1st order diffracted beam in vertical center of lens.

- Adjust M4 to place beam on vertical axis of grating, spot 2 of grating template.

- Adjust M3 vertical adjustment to place spot in corner of M5, as illustrated in photo.

- Adjust M5 to put beam on vertical axis of grating at spot 3 of grating photo and in horizontal center of M6. If necessary, use grating template.

- Measure power in and out of GDC, Should get ~50% efficiency between input port and M5.

- Replace M7

- Adjust M6 to place beam in center of M7

- Adjust M7 to center of exit port.

- Place card in GDC

- Hook up laser to fiber.

- Tighten(LIGHTLY) all three tilt/focus screws of the fiber mount to remove lens tilt.

- Use the tri-screw adjustment tool to focus/collimate beam to 1/8” dia.

- Line up two dots on card with the fiber mount with X & Y screws.

- Remove card and hook up fiber to Power Meter.

- Tweak M7 to max power.

- Tweak fiber mount X & Y screws to max power.

- Use the tri-screw adjustment tool to focus/collimate beam to max power.

- Repeat steps 19-21 until no change.

- Walk the beam with M6 and M7 to max power.

- Repeat steps 19-23 until no change.

GDC Alignment procedure

M2

Telescope Mirror

M3

First pass Mirror

M6

First output mirror

L Fourier transform lens

M4 Folding Mirror

G

Grating

1200 l/mm?

M1

Telescope Mirror

M5

Third pass mirror

M7

Second output mirror

Output port

Input port

Photo 1:  System layout

Place card here for step 14

1

2

3

4

Photo 2:  Grating Reflections

M3

M5

Photo 3: Spot locations on M3 and M5

Grating Template

Caution: Do Not, Under Any Circumstances, Touch Grating

It Can Not Be Cleaned!

Caution: Use low power to avoid burning paper.

Actual size, Cut out on doted line and use as a template.

0.958”

1-M3

2-M4

3-M5

4

Handle

Caution: do not, touch grating, it can not be cleaned. Use low power to avoid burning paper.